# HIV-1 and HIV-2 WGS ONT Protocol

Oregon State Public Health Laboratory

In Collaboration with the Association of Public Health Laboratories (APHL)

Noah C. Hull1, Eugene Yeboah1, Laura Tsaknaridis1, Vanda Makris1, John Fontana1, Shane Sevey1

Table of Contents

I.	PURPOSE:	2

II.	SCOPE / SPECIMEN ACCEPTANCE CRITERIA:	2

III.	DEFINITIONS:	2

IV.	ROLES and RESPONSIBILITIES:	3

V.	EQUIPMENT / SUPPLIES / REAGENTS:	3

VI.	Hazardous substances / Safety Precautions:	4

VII.	Hazardous chemical waste disposal:	5

VIII.	Instrument start-up / maintenance:	5

IX.	CALIBRATION:	5

X.	QUALITY CONTROL:	6

XI.	PROCEDURE:	6

XII.	QC ACCEPTANCE CRITERIA:	15

XIII.	rESULT REPORTING AND INTERPRETATION:	15

XIV.	TROUBLESHOOTING:	15

XV.	REFERENCES:	15

# PURPOSE:

Human immunodeficiency virus (HIV) is the pathogen responsible for acquired immuno-deficiency syndrome (AIDS) and continues to be a major global public health issue.  According to the World Health Organization as of 2020 there are approximately 37.7 million people living with HIV worldwide and during this year around 680,000 have died. HIV can be managed with antiretroviral (ART) drug treatments by suppressing the viral replication of HIV within the infected individual. Unfortunately, due to the nature of the virus and the availability to obtain consistent ART treatments, drug resistance can occur making the ART drug no longer effective. HIV drug resistance testing is recommended when beginning ART treatments and for evaluating treatment failure.  Traditionally HIV drug resistance testing is performed by Sanger sequencing of specific regions within the HIV genome to identify drug resistance mutations. Recently, Next Generation Sequencing (NGS) technologies have been implemented to gain better sensitivity and reproducibility when compared to sanger sequencing.  NGS for HIV resistance testing has typically been performed on both the Illumina MiSeq and Life Technologies Ion Torrent platforms.  These platforms can be costly, time consuming and provide limited availability for testing in more rural locations.  Therefore, there is clearly a place for portable NGS instruments that can be deployed into the field for HIV drug resistant testing. Oxford Nanopore Technologies (ONT) provide portable sequencers that identify DNA bases by measuring the changes in electrical conductivity generated as DNA strands pass through a biological pore. OSPHL in collaboration with APHL, will evaluate the performance of the Oxford Nanopore Technology sequencers for HIV genome sequencing using a modified protocol of the ARTIC amplicon v3.

The modifications include:

- Switching to the Nanopore RAPID barcoding, which requires less time (10 min) and fewer reagents.

- HIV Primer sequences designed by Primal Scheme. See Appendix 1.

HIV-1:

https://www.dropbox.com/sh/xnhpektoen7chdd/AAAT789gnEnwbbCjF_EMuRu8a?dl=0

HIV-2:

https://www.dropbox.com/sh/2sco4ffr30ids7a/AAD0z-VXSXElamsuys7zsmpLa?dl=0

# SCOPE / SPECIMEN ACCEPTANCE CRITERIA:

Positive HIV-1 and HIV-2 serum specimens

# DEFINITIONS:

RNA:  Ribonucleic acid; a polymeric molecule essential in various biological roles in coding, decoding, regulation and expression of genes.

RT-PCR: Reverse Transcription-PCR, a PCR technique that transcribes complimentary DNA (cDNA) from an RNA template, and then amplifies the cDNA by PCR.

PCR: Polymerase chain reaction is a process used to amplify a single or few copies of DNA.

Library: A collection of similarly sized DNA fragments with known adapter sequences added to the 5' and 3' ends.

NGS: Next-generation sequencing (NGS) is a technology for determining the sequence of DNA or RNA to study genetic variation associated with diseases or other biological phenomena.

# ROLES and RESPONSIBILITIES:

It is the responsibility of the laboratorian (Microbiologist 2 and 3) to ensure that:

- All lab policies are followed, and sterile techniques are enforced.

- All laboratory equipment is calibrated and maintained consistent with manufacturer’s recommendations.

- All reagents have not gone past their expiration date and have been stored adequately.

This procedure assumes a basic familiarity with next-generation DNA sequencing (NGS).  Laboratorians must be trained in molecular testing and the ONT Sequencing instruments prior to performing this procedure.  Refer to Good Laboratory Practices for Molecular Testing policy. (LINKED)

# EQUIPMENT / SUPPLIES / REAGENTS:

| REAGENTS |  |
| --- | --- |
| LunaScript® RT SuperMix Kit | NEB Cat# E3101 |
| Q5® Hot Start High-Fidelity SX Master Mix | NEB Cat# M0494 |
| HIV-1/HIV-2 Primers | IDT – Custom |
| Rapid Barcoding Kit: | Nanopore Cat# SQK-RBK110.96 |
| - Rapid Barcode Plate (RB96) |  |
| - SPRI beads (SPRI) |  |
| - Elution Buffer (EB) |  |
| - Rapid Adapter F (RAP-F) |  |
| - Sequencing Buffer II (SBII) |  |
| - Loading Beads II (LBII) |  |
| - Loading Solution (LS) |  |
| - Flush Tether (FLT) |  |
| - Flush Buffer (FB) |  |
| Qubit™ Fluorometry 4.0 | Cat. # Q33238 |
| Qubit™ 1X dsDNA HS Assay Kit | Cat. # Q33230 or Q33231 |
| Absolute Ethanol, 200 proof, Molecular Biology Grade | Cat. # T038181000 |
| UltraPure™ DNase/RNase-Free Distilled Water | Cat. # 10977015 |
| SpotON Flow Cell (R9.4.1) | Nanopore Cat# FLO-MIN106D |
| Flow Cell Wash Kit | Nanopore Cat# EXP-WSH004 |
| - Wash Mix (WMX) |  |
| - Diluent (DIL) |  |
| - Storage Buffer (S) |  |
| EQUIPMENT |  |
| PCR Cabinet – NuAire, model 126-300 | Cat. # A35422 |
| SimpliAmp™ Thermal Cycler or similar | Cat. # A35422 |
| Eppendorf® Centrifuge 5430/5430R | Cat. # EP022620645 |
| myFuge 12 Micro-Centrifuge | Cat. # 50-550-338 |
| Gentle rotator mixer or HULA mixer |  |
| Vortex Mixer |  |
| Rainin Pipettors – 2ml, 20ml, 200ml and 1000ml |  |
| Rainin Multichannel Pipettors – 20ml and 200ml |  |
| Rainin Repeater Pipettor – 20ml and 200ml |  |
| Eppendorf™ PCR Cooler | Cat. # 05-403-00 |
| Magnetic Separator, suitable for 1.5ml tubes |  |
| Qubit® 4.0 Fluorometer |  |
| ONT Sequencer |  |
| SUPPLIES |  |
| MicroAmp™ Optical 96-Well Reaction Plate | Cat. # N8010560 |
| MicroAmp™ Optical 8-Cap Strips | Cat. # 4323032 |
| MicroAmp™ 8-Cap Strip, clear | Cat. # N8010535 |
| Low Protein Binding Collection Tubes (1.5 ml) | Cat. # 90410 |
| Rainin 20ml filter tips, Terra rack | Cat. # 17014961 |
| Rainin 200ml filter tips, Terra rack | Cat. # 17014963 |
| Rainin 1000ml filter tips, Terra rack | Cat. # 17014967 |
| RNase Away™ | Fisher Scientific #21-236-21 or equiv. |
| Coverage Spray HB Plus | Steris #1624-77 or VWR #4212-963 |
| Disposable powder-free gloves |  |
| Permanent marking pens |  |
| Kim Wipes |  |
| Biohazard containers for sharps, assorted sizes |  |

# Hazardous substances / Safety Precautions:

BIOHAZARD

Conduct all work in properly equipped facilities with the appropriate safety equipment (for example, physical containment devices). Safety equipment can also include items for personal protection, such as gloves and coats. Individuals should be trained according to applicable regulatory and facility requirements before working with potentially biohazardous materials. Follow all applicable local, state/provincial, and/or national regulations.

- Lab coats and gloves must be worn when performing this procedure.

- The following items should be treated as biohazards after the run is complete:

  - RNA sample plates

  - RT-PCR plates

  - PCR plates

  - Reagent tubes

  - Used sample plates

  - Reagents

# Hazardous chemical waste disposal:

- Ensure use of primary and secondary waste containers. (A primary waste

container holds the immediate waste. A secondary container contains spills

or leaks from the primary container. Both containers must be compatible

with the waste material and meet federal, state, and local requirements for

container storage.)

- Ethanol waste can be disposed of in the lab sink under running water. Ensure that gloves, gowns are worn whenever using ethanol. Eye protection may be used as well.

# Instrument start-up / maintenance:

ONT Sequencer

- Hardware Check – Perform monthly

- Flowcell QC check – Perform before each run (>800 pores to pass)

SimpliAmp™ Thermal Cycler

- After each use

  - Clean the housing or external part of the thermal cycler with a soft cleaning cloth or paper with approved disinfectant.

- Annual QC Check

  - Annual QC will be done through the “self-verification test feature” on the instrument. The check includes testing the block, heated cover, and other components.

- As needed

  - The thermal area lid should be cleaned with a cloth or soft cleaning paper moistened with isopropyl alcohol.

  - Clean the wells with water followed by Isopropanol

Qubit® 4.0 Fluorometer

The Qubit® 4.0 Fluorometer does not need regular maintenance.

- Do not perform any repairs or service on the Qubit® 4.0 Fluorometer to avoid damaging the instrument.

- Do not expose the Qubit® 4.0 Fluorometer to direct sunlight.

- Clean the surface of the Qubit® 4.0 Fluorometer with a damp cloth.

- To clean the touch screen, turn off the Qubit® 4.0 Fluorometer, disconnect the power cable, and clean the touch screen with a soft cloth lightly moistened with LCD cleansing detergent.

# CALIBRATION:

SpotON Flowcell

The ONT sequencer will perform a QC check on each flowcell prior to proceeding with priming.

Thermocycler

For further information on the thermocycler system calibration and temperature maintenance procedures, refer to SimpliAmp™ Thermal Cycler User Guide (Pub. No. MAN0009889)

# QUALITY CONTROL:

- Quality controls are integrated into the sequencing assay process:

  - HIV-1 Positive Control

    - The positive sequence control should display a characteristic sequence profile and pass all quality control metrics. If the positive control fails, the sequencing run must be repeated.

  - No template control (NTC)

    - The negative sequence control (NTC) should not display a sequencing profile. If the negative control fails, repeat the sequencing run due to potential of cross contamination. Make sure to thoroughly decontaminate the sequencing area using Coverage Spray HB Plus, RNAse Away® or similar disinfectants.

  - Sequencing metrics of quality controls are analyzed, documented, and monitored.

- Always wear a lab coat and gloves. Change gloves often.

- Use dedicated pipettes and tips.

- Clean work surfaces, pipettes, and other equipment with RNase Away® prior to use.

- New lots of reagents are verified and documented.

- Assay performance metrics verified after instrument maintenance, repair, or software update.

- Take precautions to prevent cross contamination between samples.

  - An in-house carryover evaluation must be performed biannually.

# PROCEDURE:

## Reverse Transcription

The LunaScript® RT SuperMix Kit will be used to reverse transcribe HIV RNA for preparation to perform the Primer Pool PCR assay.

NOTE: LunaScript® RT SuperMix will be kept in the PCR Clean room.

- Place 96-well PCR plate on an Eppendorf® PCR cooler (stored at -20oC).

- Using a repeater pipettor aliquot 4 μl of LunaScript® RT SuperMix into each sample well of a 96‑well PCR plate according to plate map.

- Cover the plate and move the PCR plate containing master mix from the PCR Clean Room to the processing bench for sample addition.

- Add 16 μl of RNA extract to the appropriate sample well and gently mix by pipetting up and down. If less than 16 μl is available, make up the volume with nuclease-free water. Total volume = 20 μl per well.

- Seal the plate with MicroAmp™ 8-cap strips and briefly centrifuge to collect droplets.

- Return the RT plate to the Eppendorf® PCR cooler until thermocycler reaches 25oC.

- Load the plate into the SimpliAmp™ thermocycler, then run the “HIV REVERSE TRANSCRIPTION” method with the following parameters:

| REPS | 1 | 1 | 1 | Hold |
| --- | --- | --- | --- | --- |
| TIME | 2:00 | 10:00 | 1:00 | ∞ |
| TEMP | 25oC | 55oC | 95oC | 4oC |

    - Set heated lid to 105oC

    - Runtime ~14 minutes

## PCR Primer Pool Preparation

- Combine the following components to prepare the master mix for each primer pool (A or B). (Allow for pipetting loss by adding ~ 4 reactions per 32 samples.)

NOTE: Use the HIV-1 RT and Primer Pool Preparation Worksheet to calculate the required amount of master mix. This can be made at the same time as the LunaScript® SuperMix RT plate and stored at 4oC until use.

@10 μM

Component 						Volume

Nuclease-free Water 		 			8.90 μl

HIV-1 Primer Pool (A or B)				1.10 μl

Q5® Hot Start HF 2X Master Mix 			12.5 μl

Total volume per sample				22.5 μl

- Place 96-well PCR plate on an Eppendorf® PCR cooler (stored at -20oC).

- Using a repeater pipettor aliquot 22.5 μl of each master mix into the appropriate wells of the 96-well plate.

- Cover and store at 4oC until reverse transcription protocol is complete.

- Remove the RT plate from the thermocycler and centrifuge.

- Carefully remove the 8-cap strips and using a multi-channel pipettor transfer 2.5 μl of the RT reactions into the appropriate wells of the PCR plate containing the primer pools. Gently mix by pipetting the contents of each well up and down. Total volume = 25 μl per well.

NOTE: It is useful to separate the plate in half with primer pool A in columns 1-6 and primer pool B in columns 7-12.

- Seal the plate with 8-cap strips and briefly centrifuge.

- Load the plate into the SimpliAmp™ thermocycler, then run the “HIV PCR PRIMERS” method with the following parameters:

| REPS | 1 | 45X |  | Hold |
| --- | --- | --- | --- | --- |
| TIME | 0:30 | 0:15 | 5:00 | ∞ |
| TEMP | 98oC | 98oC | 65oC | 4oC |

    - Set heated lid to 105oC

    - Runtime ~5 Hours

- Repeat the same primer pool preparation from above (Step 1) and aliquot it into the same wells as in Step 1 into a new PCR plate in the same wells as the previous plate.

NOTE: Use the HIV-1 RT and Primer Pool Preparation Worksheet to calculate the required amount of master mix.

@10 μM

Component 						Volume

Nuclease-free Water 		 			8.90 μl

HIV-1 Primer Pool (A or B)				1.10 μl

Q5® Hot Start HF 2X Master Mix 			12.5 μl

Total volume per sample				22.5 μl

- Remove the cDNA plate from the thermocycler and centrifuge.

- Carefully remove the 8-cap strips and using a multi-channel pipettor transfer 2.5 μl of the 1st cDNA products into the corresponding wells of the new PCR plate containing the 2nd set of primer pools.

(2.5 μl of 1st Pool A or B goes into corresponding 2nd Pool A or B of new plate.)

Gently mix by pipetting the contents of each well up and down. Total volume = 25 μl per well.

- Seal the plate with 8-cap strips and briefly centrifuge.

- Load the plate into the SimpliAmp™ thermocycler, then run the “HIV PCR PRIMERS” method again. Same parameters as above.

| REPS | 1 | 45X |  | Hold |
| --- | --- | --- | --- | --- |
| TIME | 0:30 | 0:15 | 5:00 | ∞ |
| TEMP | 98oC | 98oC | 65oC | 4oC |

    - Set heated lid to 105oC

    - Runtime ~5 Hours

## Addition of rapid barcodes

- Thaw Rapid Barcode Plate and bring SPRI beads to room temperature ~ 1 hour before PCR is complete.

- Spin down the Rapid Barcode Plate.

- Remove the primer-specific cDNA plate from the thermocycler and centrifuge briefly to collect the contents at the bottom of the wells.

- Place the plate on an Eppendorf® PCR cooler.

- Carefully remove the 8-cap strips from the cDNA plate to prevent any splashing.

- Place a clean 96-well PCR plate on a separate Eppendorf® PCR cooler (stored at       -20oC) and combine each 25 μl cDNA reaction into a single well in the plate. Mix by pipetting up and down and set aside. Total volume per sample = 50 μl.

- Place another clean PCR 96-well sample plate on a separate Eppendorf® PCR cooler and using a repeater pipettor aliquot 2.5 μl of nuclease-free water into corresponding sample wells.  This is the Barcode Attachment Plate (BAP).

- Using a multichannel pipette, transfer 5 μl of the pooled cDNA products to the corresponding well of the Barcode Attachment Plate (BAP) and mix by pipetting.

- Using a multichannel pipette, transfer 2.5 μl of the appropriate Rapid Barcodes to the corresponding well of the Barcode Attachment Plate (BAP) and mix by pipetting. Be careful not to cross-contaminate the different wells. Total volume per sample = 10 μl.

- Seal the plate with 8-cap strips and briefly centrifuge.

- Incubate the plate on the thermocycler using the “RAPID BARCODE ATTACHMENT” protocol with the following parameters:

| REPS | 1 | 1 |
| --- | --- | --- |
| TIME | 2:00 | 2:00 |
| TEMP | 30oC | 80oC |

## Pooling Samples and Clean-up

- Remove the Barcode Attachment Plate (BAP) from the thermocycler and centrifuge briefly to collect the contents at the bottom of the wells.

- Carefully remove the 8-cap strips from the BAP plate to prevent any splashing.

- Pool the barcoded samples into a 1.5 ml Eppendorf DNA LoBind tube. ~10 μl per sample

For example:

| # Samples | 24 | 48 | 96 |
| --- | --- | --- | --- |
| Total Volume | ~240 μl | ~480 μl | ~960 μl |

- Resuspend the SPRI beads by vortexing.

- Add an equal volume of resuspended SPRI beads to the pooled samples and mix by flicking the tube.

For example:

| # Samples | 24 | 48 | 96 |
| --- | --- | --- | --- |
| Volume of SRPI beads | ~240 μl | ~480 μl | ~960 μl |

- Incubate at room temperature on a gentle rotator mixer (~700 rpm) for 5 minutes.

- While incubating, prepare 4 ml of fresh 80% ethanol in nuclease-free water.

- Spin down the sample tube briefly to bring any liquid down from the cap and place it on the magnet. A pellet will form on the side wall of the tube.

- When the solution clears, keep the tube on the magnet and pipette off the supernatant and discard.

- Without disturbing the beads, add 1.5 ml of the freshly prepared 80% ethanol. Rotate the tube on the magnet to move the beads front to back a few times to wash beads. Remove the ethanol using a pipette and discard.

- Repeat the previous step.

- Briefly spin down the tube and place it back on the magnet. Pipette off any residual ethanol. Allow to dry for 30 seconds. Do not over dry the pellet to the point of cracking.

- Remove the tube from the magnet and resuspend the pellet in 30 μl Elution Buffer (EB). Gently mix by pipetting up and down.

- Incubate for 10 minutes at room temperature.

- Place the tube on the magnet until the eluate is clear and colorless.

- Remove and transfer the 30 μl eluate (containing the DNA library) into a clean 1.5 ml Eppendorf DNA LoBind tube. Dispose of the pelleted beads. Be careful not to transfer any of the beads.

- Proceed to quantifying the DNA concentration using the Qubit dsDNA HS Assy Kit.

## Library quantification and normalization

Analyze 1 μl of the amplified library using the Qubit™ 4.0 Fluorometer and the Qubit™ dsDNA HS Assay Kit. For more information, see the Qubit™ dsDNA HS Assay Kits User Guide.

      - Determine the amplified library concentration using Qubit™ 4.0 Fluorometer:

- Prepare the Qubit™ standards as directed in the user guide (10 μl standard  + 190 μl Qubit™ dsDNA HS Buffer ), mix well, and incubate for at least 2 minutes.

- For each sample, combine 1 μl of the amplified library with 199 μl of Qubit™ dsDNA HS Buffer, mix well, and incubate for at least 2 minutes.

- On the Qubit™ 4.0 Fluorometer home screen, select dsDNA then 1x dsDNA High Sensitivity.

- Measure the prepared standards.

- Select the sample concentration as “ng/μl” with an input value of 1 μl.

- Measure and record the library concentration.

      - In a new LoBind tube normalize a library pool to 600-800 ng with EB. A total volume of 11 μl is required for the following step.

Calculate appropriate dilution volumes using the following:

Volume Library Pool (μl)  =  ___Target quantity (ng)___

Pool Concentration (ng/μl)

Volume EB (μl) = 11 μl – Volume Library Pool (μl)

If concentration of the total pool is less than 600ng/μl, use library pool undiluted.

      - Add 1 μl of Rapid Adapter F (RAP F) to the 11 μl normalized barcoded DNA and mix gently.

      - Incubate at room temperature for 5 minutes.

After incubation, place the prepared library in a cold rack until ready to load onto the SpotON flowcell.

## Priming and Loading the SpotON Flow Cell

- Thaw the Sequencing Buffer II (SBII), Loading Beads II (LBII) or Loading Solution (LS, if using), Flush Tether (FLT) and Flush Buffer (FB) at room temperature.

- Mix the SBII, FB and FLT tubes by vortexing. Spin down the SBII and FLT tubes.

- Open the ONT sequencer lid and slide the flow cell under the clip. Press down firmly on the flow cell to ensure correct thermal and electrical contact.

- QC the SpotON flowcell using the MinKNOW software before proceeding. QC passes if total active pores are >800.

- After QC is complete, move the unit including the flowcell to the bench for priming and loading.

- Slide the priming port cover clockwise to open the priming port.

- Prime and load the SpotON Flow Cell

  - In a 1.5 ml Eppendorf LoBind tube, prepare the flow cell priming mix by adding 30 μl of Flush Tether (FLT) to 1.17 ml of Flush Buffer (FB) and vortex to mix.

  - Open the priming port and check for a small air bubble under the cover. Draw back a small volume to remove any bubbles (a few μl):

    - Set a P1000 to 200 μl

    - Insert the tip into the priming port

    - Turn the wheel until the dial shows 220-230 μl, or until you can see a small volume of buffer entering the pipette tip.

    - Visually inspect that there is continuous buffer from the priming port across the sensor array.

  - Load 800 ul of the priming mix into the flow cell via the priming port without introducing bubbles. Wait 5 minutes.

- Prepare the library for loading

  - Thoroughly mix the contents of the Loading Beads II (LBII) by pipetting up and down. The LBII tube contains a suspension of beads that settle very quickly. It is vital that they are mixed immediately before use.

  - In a new tube, prepare the library for loading as follows:

Reagent					Volume

Sequencing Buff II (SBII) 			37.5 μl

Loading Beads II				22.5 μl

DNA Library			 		12 μl

Total volume					75 μl

- Complete the flow cell priming

  - Gently lift the SpotON sample port cover to make the sample port accessible.

  - Slowly load 200 ul of the priming mix into the flow cell via the priming port (not the SpotON sample port), avoiding the instruction of bubbles.

  - Immediately, mix the prepared library by pipetting up and down gently.

  - Add 75 μl of sample to the flow cell via the SpotON port in a dropwise manner. Ensure each drop flows into the port before adding the next.

  - Replace the SpotON sample port cover, making sure it is seated correctly, close the priming port, and close the sequencer lid.

## Starting and monitoring the sequencing run using the MinKNOW software

## Washing the Flowcell  – Perform immediately after sequencing

- Vortex Wash Mix (WMX) and Storage Buffer (S).

- Make up the Wash Buffer Solution (Mix 20 μl Wash Mix (WMX) and 380 μl Diluent (DIL)).

- Remove the flow cell from the instrument and place it in the packaging without the cover.

- Open the priming port.

- Slowly remove all the waste from the waste port and discard.

- Slowly add 400 μl wash solution to the priming port.

- Incubate at room temperature for at least 10 minutes (ideally 1 hour).

- Add 500 μl Storage Buffer (S) into the priming port and close the port.

- Store at 4-8oC

You must QC the flowcell again before use to ensure a suitable number of active pores.

# QC ACCEPTANCE CRITERIA:

The positive sequence control should display a characteristic sequence profile and pass all quality control metrics below. If the positive control fails, the sequencing run must be repeated.

The negative sequence control (NTC) should not display a sequencing profile. If the negative control fails, repeat the sequencing run due to potential of cross contamination.

| Quality Control Metrics | Acceptable Range |
| --- | --- |
| Target base coverage 20X |  |
| Coverage Depth |  |
| Percent N's |  |
| Total Length |  |

# rESULT REPORTING AND INTERPRETATION:

N/A

# TROUBLESHOOTING:

# REFERENCES:

- HIV-1 and HIV-2 genome sequencing protocol (400 bp amplicon primer set, using Nanopore Rapid kit); Noah C. Hull, Eugene Yeboah, APHL 2022

- MinION flow cell wash protocol; Dr Linzy Elton, University College London, UK, Alp Aydin, Quadram Institute, UK, Dewi Owen, University College London, UK, Eloise Rose, University College London, UK

Appendix 1: Primer Sequences

HIV-1:

| name | pool | seq | size | %gc | tm (use 65) |
| --- | --- | --- | --- | --- | --- |
| HIV-1_v1.0_1_LEFT | 1 | TGGTTAGACCAGATCTGAGCCT | 22 | 50 | 60.48 |
| HIV-1_v1.0_1_RIGHT | 1 | TTTCTTTCCCCCTGGCCTTAAC | 22 | 50 | 60.68 |
| HIV-1_v1.0_3_LEFT | 1 | GCTTTAGACAAGATAGAGGAAGAGCA | 26 | 42.31 | 60.85 |
| HIV-1_v1.0_3_RIGHT | 1 | TTCCTGCTATGTCACTTCCCCT | 22 | 50 | 61.02 |
| HIV-1_v1.0_5_LEFT | 1 | TTGGATGACAGAAACCTTGTTGG | 23 | 43.48 | 59.62 |
| HIV-1_v1.0_5_RIGHT | 1 | AAGAAAATTCCCTGGCCTTCCC | 22 | 50 | 61.01 |
| HIV-1_v1.0_7_LEFT | 1 | ATTAGAAGAAATGAGTTTGCCAGGAA | 26 | 34.62 | 59.72 |
| HIV-1_v1.0_7_RIGHT | 1 | TTCTTTATGGCAAATACTGGAGTATTGT | 28 | 32.14 | 59.98 |
| HIV-1_v1.0_9_LEFT | 1 | GAGACACCAGGGATTAGATATCAGT | 25 | 44 | 59.61 |
| HIV-1_v1.0_9_RIGHT | 1 | CCCTGGGTAAATCTGACTTGCC | 22 | 54.55 | 61.13 |
| HIV-1_v1.0_11_LEFT | 1 | AGAGCCATTTAAAAATCTGAAAACAGGA | 28 | 32.14 | 60.77 |
| HIV-1_v1.0_11_RIGHT | 1 | CAGTCTTCTGATTTGTTGTGTCAGT | 25 | 40 | 60.08 |
| HIV-1_v1.0_13_LEFT | 1 | GTCAGTGCTGGAATCAGGAAAGT | 23 | 47.83 | 61.06 |
| HIV-1_v1.0_13_RIGHT | 1 | CGTAGCACCGGTGAAATTGCT | 21 | 52.38 | 61.64 |
| HIV-1_v1.0_15_LEFT | 1 | AGACATAATAGCAACAGACATACAAACT | 28 | 32.14 | 59.77 |
| HIV-1_v1.0_15_RIGHT | 1 | CCAATCTAGCATCCCCTAGTGG | 22 | 54.55 | 59.88 |
| HIV-1_v1.0_17_LEFT | 1 | CAAGCAGGACATAACAAGGTAGGA | 24 | 45.83 | 60.65 |
| HIV-1_v1.0_17_RIGHT | 1 | TCCAGGGCTCTAGTCTAGGATC | 22 | 54.55 | 59.81 |
| HIV-1_v1.0_19_LEFT | 1 | TCTCTATCAAAGCAGTAAGTAGTACATGT | 29 | 34.48 | 60.52 |
| HIV-1_v1.0_19_RIGHT | 1 | GCATGTGTGGCCCAAACATTAT | 22 | 45.45 | 60.28 |
| HIV-1_v1.0_21_LEFT | 1 | AGCGGGAGAATGATAATGGAGAA | 23 | 43.48 | 59.61 |
| HIV-1_v1.0_21_RIGHT | 1 | GCATTGTCCGTGAAATTGACAGA | 23 | 43.48 | 60.31 |
| HIV-1_v1.0_23_LEFT | 1 | AGCTAGCAAATTAAGAGAACAATTTGGA | 28 | 32.14 | 60.61 |
| HIV-1_v1.0_23_RIGHT | 1 | TTCACTTCTCCAATTGTCCCTCA | 23 | 43.48 | 59.86 |
| HIV-1_v1.0_25_LEFT | 1 | CTATTGAGGCGCAACAGCATCT | 22 | 50 | 61.5 |
| HIV-1_v1.0_25_RIGHT | 1 | ACCTACCAAGCCTCCTACTATCA | 23 | 47.83 | 60.05 |
| HIV-1_v1.0_27_LEFT | 1 | ACCACCGCTTGAGAGACTTACT | 22 | 50 | 61.27 |
| HIV-1_v1.0_27_RIGHT | 1 | TGCTCCATGTTTTTCCAGGTCT | 22 | 45.45 | 60.34 |
| HIV-1_v1.0_29_LEFT | 1 | CACACACAAGGCTACTTCCCTG | 22 | 54.55 | 61.31 |
| HIV-1_v1.0_29_RIGHT | 1 | AACCAGAGAGACCCAGTACAGG | 22 | 54.55 | 61.01 |
| HIV-1_v1.0_2_LEFT | 2 | TAGAAGGAGAGAGATGGGTGCG | 22 | 54.55 | 61.19 |
| HIV-1_v1.0_2_RIGHT | 2 | TTTTGGCTGACCTGATTGCTGT | 22 | 45.45 | 61.2 |
| HIV-1_v1.0_4_LEFT | 2 | GCTGCAGAATGGGATAGAGTGC | 22 | 54.55 | 61.56 |
| HIV-1_v1.0_4_RIGHT | 2 | TTCTTCTAGTGTAGCCGCTGGT | 22 | 50 | 61.33 |
| HIV-1_v1.0_6_LEFT | 2 | GGAAGGACACCAAATGAAAGATTGT | 25 | 40 | 60.37 |
| HIV-1_v1.0_6_RIGHT | 2 | TGTCCACAGATTTCTATGAGTATCTGA | 27 | 37.04 | 59.94 |
| HIV-1_v1.0_8_LEFT | 2 | AGTAGAAATTTGTACAGAGATGGAAAAGG | 29 | 34.48 | 60.47 |
| HIV-1_v1.0_8_RIGHT | 2 | AAGGCTCTAAGATTTTTGTCATGCT | 25 | 36 | 59.67 |
| HIV-1_v1.0_10_LEFT | 2 | CAGCCTATAGTGCTGCCAGAAA | 22 | 50 | 60.6 |
| HIV-1_v1.0_10_RIGHT | 2 | TTTGCACTGCCTCTGTTAATTGT | 23 | 39.13 | 59.62 |
| HIV-1_v1.0_12_LEFT | 2 | GGGAGACTAAATTAGGAAAAGCAGGA | 26 | 42.31 | 61.02 |
| HIV-1_v1.0_12_RIGHT | 2 | AGCCATTGCTCTCCAATTACTGT | 23 | 43.48 | 60.57 |
| HIV-1_v1.0_14_LEFT | 2 | GGGCAGGAAACAGCATATTTTCT | 23 | 43.48 | 59.81 |
| HIV-1_v1.0_14_RIGHT | 2 | TGCTGTCCCTGTAATAAACCCG | 22 | 50 | 60.54 |
| HIV-1_v1.0_16_LEFT | 2 | GGGAAAGCTAGGGGATGGTTTT | 22 | 50 | 60.48 |
| HIV-1_v1.0_16_RIGHT | 2 | TCGTAACACTAGGCAAAGGTGG | 22 | 50 | 60.47 |
| HIV-1_v1.0_18_LEFT | 2 | GCAACAACTGCTGTTTATCCATTTT | 25 | 36 | 59.91 |
| HIV-1_v1.0_18_RIGHT | 2 | TTTCCTATATTCTATGATTACTATGGACCAC | 31 | 32.26 | 59.61 |
| HIV-1_v1.0_20_LEFT | 2 | TACCTGTGTGGAAGGAAGCAAC | 22 | 50 | 60.67 |
| HIV-1_v1.0_20_RIGHT | 2 | TGCATATTCTTTCTGCACCTTACCT | 25 | 40 | 60.96 |
| HIV-1_v1.0_22_LEFT | 2 | GCCAGTAGTATCAACTCAACTGCT | 24 | 45.83 | 60.94 |
| HIV-1_v1.0_22_RIGHT | 2 | ACAGTAGAAAAATTCCCCTCCACA | 24 | 41.67 | 60.22 |
| HIV-1_v1.0_24_LEFT | 2 | GGGCTGCTATTAACAAGAGATGGT | 24 | 45.83 | 61.01 |
| HIV-1_v1.0_24_RIGHT | 2 | AGGTATCTTTCCACAGCCAGGA | 22 | 50 | 61.02 |
| HIV-1_v1.0_26_LEFT | 2 | TGGGCAAGTTTGTGGAATTGGT | 22 | 45.45 | 61.41 |
| HIV-1_v1.0_26_RIGHT | 2 | ACCAATATTTGAGGGCTTCCCAC | 23 | 47.83 | 61.14 |
| HIV-1_v1.0_28_LEFT | 2 | TGGATGGCCTACTGTAAGGGAA | 22 | 50 | 60.75 |
| HIV-1_v1.0_28_RIGHT | 2 | AGCTTGTAGCACCATCCAAAGG | 22 | 50 | 61.06 |

HIV-2:

| name | pool | seq | size | %gc | tm (use 65) |
| --- | --- | --- | --- | --- | --- |
| HIV-2_v1.0_1_LEFT | 1 | TGCAAGGGATGTTTTACAGTAGGA | 24 | 41.67 | 60.28 |
| HIV-2_v1.0_1_RIGHT | 1 | CCAAGTATGGTTGTTCCTGTTATTCA | 26 | 38.46 | 60.01 |
| HIV-2_v1.0_3_LEFT | 1 | GCATTGTATTCAGTCGCTCTGC | 22 | 50 | 60.46 |
| HIV-2_v1.0_3_RIGHT | 1 | ACTCCGTCGTGGTTTGTTCCT | 21 | 52.38 | 62.01 |
| HIV-2_v1.0_5_LEFT | 1 | AAAACATATTGTGTGGGCAGCG | 22 | 45.45 | 60.53 |
| HIV-2_v1.0_5_RIGHT | 1 | TCCTCCACTAATTTTACCCATGCAT | 25 | 40 | 60.73 |
| HIV-2_v1.0_7_LEFT | 1 | ATTCGCAGCACCCAATACCAG | 21 | 52.38 | 61.12 |
| HIV-2_v1.0_7_RIGHT | 1 | GCGGTTAGCATCTCTTCTAGGG | 22 | 54.55 | 60.47 |
| HIV-2_v1.0_9_LEFT | 1 | CAGGACACATCATGGCAAACTG | 22 | 50 | 60.27 |
| HIV-2_v1.0_9_RIGHT | 1 | TTGTCCCCTAATTCTATTCCTGCT | 24 | 41.67 | 59.6 |
| HIV-2_v1.0_11_LEFT | 1 | GGAAAGATGGACCAAGGCTGAA | 22 | 50 | 60.74 |
| HIV-2_v1.0_11_RIGHT | 1 | CTTCCATCCTTGTGGCAAGACT | 22 | 50 | 60.74 |
| HIV-2_v1.0_13_LEFT | 1 | GGGTTTTCTACTCCAGATGAGAAGT | 25 | 44 | 60.43 |
| HIV-2_v1.0_13_RIGHT | 1 | ACTTTTAGGATTTTCTCTTCCTGGTGT | 27 | 37.04 | 61.14 |
| HIV-2_v1.0_15_LEFT | 1 | GGTTAACATTTAACCTAGTAGGAGATCCT | 29 | 37.93 | 60.78 |
| HIV-2_v1.0_15_RIGHT | 1 | ACTTGTCTGATGCCTTGACTTACT | 24 | 41.67 | 60.28 |
| HIV-2_v1.0_17_LEFT | 1 | TCATTATAGTAGCAGTACATGTTGCAA | 27 | 33.33 | 59.51 |
| HIV-2_v1.0_17_RIGHT | 1 | CATGTTGATTAGTCTTTCTGCTGGG | 25 | 44 | 60.54 |
| HIV-2_v1.0_19_LEFT | 1 | ACTGGATAGTAGTCCCCACCTG | 22 | 54.55 | 60.55 |
| HIV-2_v1.0_19_RIGHT | 1 | ATGAGCTTGGGGATAGTTGCAG | 22 | 50 | 60.61 |
| HIV-2_v1.0_21_LEFT | 1 | GGGATGTCAGCAAGCTACACAA | 22 | 50 | 61.06 |
| HIV-2_v1.0_21_RIGHT | 1 | GGTTACATCCCGCTCTGAAGTG | 22 | 54.55 | 61.17 |
| HIV-2_v1.0_23_LEFT | 1 | TGGTAGGAATCAGCTGTTTGTTGT | 24 | 41.67 | 61.01 |
| HIV-2_v1.0_23_RIGHT | 1 | ATGTGGGTGGTCTCGAGGTT | 20 | 55 | 60.85 |
| HIV-2_v1.0_25_LEFT | 1 | AGCACTATTGGGATGCTATGAGG | 23 | 47.83 | 60.25 |
| HIV-2_v1.0_25_RIGHT | 1 | CTTCCCTCCATCTGCCTCCAAA | 22 | 54.55 | 62.21 |
| HIV-2_v1.0_27_LEFT | 1 | AGCAGATAATTAATACCTGGCATAAAGT | 28 | 32.14 | 59.66 |
| HIV-2_v1.0_27_RIGHT | 1 | CTGTTGCTGTTGCTGCACTATC | 22 | 50 | 60.59 |
| HIV-2_v1.0_29_LEFT | 1 | AATATGACATGGCAGGAGTGGG | 22 | 50 | 60.34 |
| HIV-2_v1.0_29_RIGHT | 1 | CAAGGCCACAAGTCGTAACCA | 21 | 52.38 | 60.91 |
| HIV-2_v1.0_31_LEFT | 1 | ATTCGCGAGGACTACGAGAGAG | 22 | 54.55 | 61.55 |
| HIV-2_v1.0_31_RIGHT | 1 | CCCTTCCTCTTTTTCTAGGTATATGTCT | 28 | 39.29 | 60.51 |
| HIV-2_v1.0_33_LEFT | 1 | ACTACAAGGCCTTCACTCTGTAC | 23 | 47.83 | 60 |
| HIV-2_v1.0_33_RIGHT | 1 | AGTACCGGCCAAGTACTGGT | 20 | 55 | 60.56 |
| HIV-2_v1.0_2_LEFT | 2 | GAGTTTGGGCACAAGTCAGGAT | 22 | 50 | 61 |
| HIV-2_v1.0_2_RIGHT | 2 | ACTGGCAGCTTTATTGAAGAGGT | 23 | 43.48 | 60.5 |
| HIV-2_v1.0_4_LEFT | 2 | CGAACAGGGACTTGAAGAGGAC | 22 | 54.55 | 60.79 |
| HIV-2_v1.0_4_RIGHT | 2 | TGGCAACCTTCTTTTGACTCCA | 22 | 45.45 | 60.54 |
| HIV-2_v1.0_6_LEFT | 2 | AAAAGAGGAAACTACCCCGTGC | 22 | 50 | 60.99 |
| HIV-2_v1.0_6_RIGHT | 2 | TCTGTAGATGTTCCCTACCGGT | 22 | 50 | 60.21 |
| HIV-2_v1.0_8_LEFT | 2 | TGGATGACCCAAACGCTGCTA | 21 | 52.38 | 62.16 |
| HIV-2_v1.0_8_RIGHT | 2 | GGGAAGTTGCGAGGCTTCTTTC | 22 | 54.55 | 62.2 |
| HIV-2_v1.0_10_LEFT | 2 | GTCACAGCGTACATCGAGGATC | 22 | 54.55 | 61.03 |
| HIV-2_v1.0_10_RIGHT | 2 | AGCTGGCCCTCTTTTTCCATTT | 22 | 45.45 | 60.95 |
| HIV-2_v1.0_12_LEFT | 2 | CTGCATTTACCCTACCAGCAGT | 22 | 50 | 60.54 |
| HIV-2_v1.0_12_RIGHT | 2 | GCTTTTGGATGTCATTGACTGTCC | 24 | 45.83 | 60.93 |
| HIV-2_v1.0_14_LEFT | 2 | ACCAAGAAGAAAAAGAATTAGAGGCA | 26 | 34.62 | 59.67 |
| HIV-2_v1.0_14_RIGHT | 2 | TCCTGCTTTTCCCTCTTTTGACT | 23 | 43.48 | 60.38 |
| HIV-2_v1.0_16_LEFT | 2 | AGCAGTCTATGTTGCATGGGTC | 22 | 50 | 60.86 |
| HIV-2_v1.0_16_RIGHT | 2 | GAAGAGTGCTGTCTGCCTTCCT | 22 | 54.55 | 62.44 |
| HIV-2_v1.0_18_LEFT | 2 | ACAATAGAAACAATAGTACTGATGGCAG | 28 | 35.71 | 60.29 |
| HIV-2_v1.0_18_RIGHT | 2 | CCTCCTCTAGGTCTTTTGTTCTGT | 24 | 45.83 | 60.04 |
| HIV-2_v1.0_20_LEFT | 2 | TCTTGCTTTACGGCAGGTGAAG | 22 | 50 | 61.31 |
| HIV-2_v1.0_20_RIGHT | 2 | GCAAGTGCACCCTCTCTTGAAA | 22 | 50 | 61.51 |
| HIV-2_v1.0_22_LEFT | 2 | GACATGGAGACACCCTTGAAGG | 22 | 54.55 | 60.8 |
| HIV-2_v1.0_22_RIGHT | 2 | CACGCGGGTATGCCATAGAAAA | 22 | 50 | 61.5 |
| HIV-2_v1.0_24_LEFT | 2 | TGGCAATGAATTGTAGCAGGGT | 22 | 45.45 | 60.74 |
| HIV-2_v1.0_24_RIGHT | 2 | AGCCTGAATAGTTGGTATCATTACATCT | 28 | 35.71 | 60.77 |
| HIV-2_v1.0_26_LEFT | 2 | TGTCAGGATTAGTGTTTCACTCTCA | 25 | 40 | 60.08 |
| HIV-2_v1.0_26_RIGHT | 2 | GCTGGTTACTGTTGATTCACAGG | 23 | 47.83 | 60.06 |
| HIV-2_v1.0_28_LEFT | 2 | CGTGCTAGGGTTCTTGGGTTTT | 22 | 50 | 61.26 |
| HIV-2_v1.0_28_RIGHT | 2 | TTCTTGTTGAATTTGGGCTTCTTCT | 25 | 36 | 60.02 |
| HIV-2_v1.0_30_LEFT | 2 | CCCGGTTATCTCCAACAGATCC | 22 | 54.55 | 60.41 |
| HIV-2_v1.0_30_RIGHT | 2 | CTTGGAACTGCGAGTATTCCCC | 22 | 54.55 | 61.18 |
| HIV-2_v1.0_32_LEFT | 2 | TCACATTTTATAAAAGAAAAAGGGGGACT | 29 | 31.03 | 60.58 |
| HIV-2_v1.0_32_RIGHT | 2 | CCCTCTTGCTTTCAGTTTTGCC | 22 | 50 | 60.73 |
| HIV-2_v1.0_34_LEFT | 2 | ACTTTCCAGAAGGGGCTGTAAC | 22 | 50 | 60.41 |
| HIV-2_v1.0_34_RIGHT | 2 | TGCTAGGGATTTTCCTGCCTTG | 22 | 50 | 60.81 |