APHL NGS Training

Eugene Yeboah1   and Noah Hull1

Association of Public Health Laboratories, Silver Spring, MD, USA1

Flow cell washing protocol (06-03-2022)

The Nanopore Flow Cell Wash kit contains:

Wash Buffer A (0.5 ml) / Wash Buffer B (1 ml) / Storage Buffer (1.6 ml)

Washing Procedure

Vortex wash buffer A and storage buffer (S)

- Remove the flow cell from the sequencing device

- Open the priming port

- Slowly remove all of the waste from the waste port (2) and discard

- Make the wash buffer solution

  - 2 ul wash buffer A

  - 398 ul wash buffer B

- Slowly add 400 µl wash solution into priming port

- Incubate at room temperature for 10 minutes

- Add 500 µl storage buffer (S) into the priming port Store at 4-8°C

- QC the flow cell and record the pore numbers. QC again before use